KIWI-OSD

Video overlay of GPS precision timestamps

Matt Montanaro

Dr. Michael Richmond

Rochester Institute of Technology

February 2005

Contents:

Purpose of KIWI OSD	 3

Contact information	 3

Parts List	 4

Circuit Diagrams	5,6

IC Pin Assignments	7

Garmin GPS 16-HVS	8

Power Supply	10

Construction	11

Using the KIWI OSD	12

INFO Switch	13

Example of KIWI OSD in action	14

Epilogue	16

Purpose of the KIWI-OSD:

The KIWI-OSD device timestamps video frames to millisecond precision. A video feed from a source (video camera) is sent into the KIWI. A timing signal from a GPS receiver is used to timestamp the incoming video with the Universal Time (UTC) signal. The video is then outputted to a receiver (TV, VCR, etc.).

Geoff Hitchcox contact information:

Geoff is the inventor of the KIWI-OSD. His can be contacted at: kiwi_36_nz@yahoo.co.nz

*** The PIC16F628-20 Integrated Circuit chip must be programmed by Geoff. He installs his KIWI OSD software on it and will mail it to you (see diagram page 5).

*** The GPS must be configured to emit the proper timing signal. Geoff has a nice program that he can email you that does this (see page 9).

Geoff’s KIWI-OSD webpages:

- http://www.geocities.com/kiwi_36_nz/kiwi_osd/kiwi_osd.htm

- http://www.geocities.com/kiwi_36_nz/kiwi_osd/example.htm

Parts List:

Name				Category		DigiKey Part#		     Quantity

LM1881 DIP-8		IC chip			LM1881N-ND			1

PIC16F628-20 DIP-20 *	IC chip			PIC16F628A-I/P-ND		1

74HC74 DIP-14		IC chip			296-1602-5-ND		1

74HC00 DIP-14		IC chip			296-12769-5-ND		1

18-pin DIP socket		DIP socket		ED58183-ND			1

14-pin DIP socket		DIP socket		ED58143-ND			2

8-pin DIP socket		DIP socket		ED58083-ND			1

680 kΩ				Resistor		680KEBK-ND			1

10 kΩ				Resistor		770-103-R10K-ND		1

22 Ω				Resistor		22H-ND			1

1 kΩ				Resistor		1.0KEBK-ND			1

1 MΩ				Resistor		1.0MEBK-ND			1

100 Ω **			Resistor		100QBK-ND			1

10 kΩ **			Resistor		10KEBK-ND			1

1 kΩ Trimpot			Potentiometer		31G13-ND			1

0.1 μF 16V			Capacitor		BC1084CT-ND		6

15 pF				Capacitor		BC1003CT-ND		2

1N914				Diode			1N914BCT-ND		2

NPN 2N2222 (TO-92)	Transistor		P2N2222AOS-ND		1

RCA video jack		Connector		CP-1403-ND			2

Power Jack 2.1mm		Connector		SC1152-ND			1

Crystal	16Mhz			Crystal			300-6034-ND			1

Push Button Switch		Switch			EG1101-ND			1

Switch cap, square		Switch			EG1192-ND			1

Other parts:

2” x 3” x 6” Plastic Case	Enclosure		Radio Shack			1

Proto-board			Prototype board	Local electronics store	1

8-pin RJ-45 jack		Connector		Local electronics store	1

9-pin DB-9 Serial		Connector		Local electronics store	1

* Must have KIWI software installed (see page 3)

** Required for Voltage Divider (see page 10)

Pin Assignments for KIWI OSD

LM1881

Pin connects to

---------------

1   pin 3 of 16F628 (Csync)

2   Composite video input via 0.1uF

3   pin 2 of 16F628 (Vsync)

4   GND

5   pin 10 of 74HC74 (PORCH) [ A ]

6   to filter using 680K and 0.1uF

7   pin 1 of 16F628 (Odd/Even)

8   +5 Volts

16F628

Pin connects to

---------------

1   RA2 to pin 7 of LM1881 (Odd/Even)

2   RA3 to pin 3 of LM1881 (Vsync)

3   RA4 to pin 1 of LM1881 (Csync)

4   MCLR to 10 K and optional RESET switch

5   GND

6   RB0 to 1PPS

7   RB1 to RS232 via transistor inverter

8   not used

9   not used

10  not used

11  not used

12  not used

13  RB7 (overlay video out)

14  +5 Volt

15  RA6 to pin 11 of 74HC74  [ C ]

16  OSC 1 CLKIN  from pin 6 74HC00 (Gated 16 MHz CLOCK)  [ B ]

17  RA0 - jumper to GND for 1PPS LO to HI UTC transition (most common)

- jumper to +5V for 1PPS Hi to Lo UTC    transition

18  RA1 connects to "info" push button switch

74HC00

Pin connects to

---------------

1   GND

2   GND

3   not used

4   to pin  9 of 74HC74

5   to pin  8 of this IC

6   to pin 16 of 16F628 [ B ]

7   GND

8   to pin  5 of this IC

9,10,11 connected together to XTAL

12  from XTAL

13  + 5 Volts

14  + 5 Volts

74HC74

Pin connects to

---------------

1  GND

2  GND

3  GND

4  GND

5  not used

6  not used

7  GND

8  not used

9  to pin 4 of 74HC00

10 to pin 5 of LM1881 (PORCH) [ A ]

11 to pin 15 of 16F628  RA6   [ C ]

12 GND

13 + 5 Volts

14 + 5 Volts

Garmin GPS 16-HVS:

Power requirements:	Voltage:	6 – 40 Vdc unregulated

Current:	100 mA @ 6 Vdc

65 mA @ 12 Vdc

28 mA @ 40 Vdc

RJ-45 Connector:

Pin 1 (Red)		- Power (+)

Pin 2 (Black)		- Ground [Power (-) and Data Return]

Pin 3 (Yellow)		- Remote Power On/Off

Pin 4 (Blue)		- Port 1 NMEA 0183

Pin 5 (White)		- Port 1 NMEA 0183

Pin 6 (Gray)		- Pulse Per Second Output

Pin 7 (Green)		- Port 2 RTCM SC-104

Pin 8 (Violet)		- Port 2 Data Output

Garmin GPS 16-HVS manual:

http://www.garmin.com/manuals/GPS17N_GPS16_17NSeriesTechnicalSpecification.pdf

Configuring the GPS:

The Garmin GPS must be configured using a PC computer before using it with the KIWI. A cable must be constructed to connect the RJ-45 connector of the GPS to the computer’s serial port. The diagram below shows the wiring for the cable:

The power source can be any DC supply between 6 and 40 volts. Once the PC and GPS are connected and powered up, the GPS can be programmed. The purpose of configuring the GPS is to enable the 1PPS signal and to only allow the NMEA sentences GGA and RMC that are needed my the KIWI. Geoff Hitchcox emailed me the program (SNSRCFG.exe) and configuration file (KIWI_OSD.cfg). (See page 3 for his contact information).

Software Settings:

- Run the SNSRCFG.exe program from Geoff

- Select “GPS 16/17”  for the Garmin Base Model

- Comm Menu > Setup → Select COM1 serial port and Auto Baud Rate

- File > Open → select KIWI_OSD.cfg → Open

- Config Menu > Sensor Config → Input your latitude and longitude

- Comm Menu > Connect

- The computer will attempt to connect to the GPS

- After successful connect: Config Menu > Send config to GPS (this will upload the .cfg file settings to the GPS)

- All done! Close program and disconnect the GPS.

Connecting the GPS to KIWI:

GPS Pins:

Pin 1 	to Power (+)

Pin 2,3 	to KIWI ground

Pin 5 	to 10k resistor “GPSTX NMEA 4800 8N1”

Pin 6	to Pin 6 of the PIC16F628-20 “GPS 1PPS” RB0

Note: Pin 17 of the PIC16F628-20 (RA0) “1PPS Edge Select” connect to ground

Power Supply:

I used a Vernier AC Adaptor Class 2 Transformer

Input: 120V AC 60Hz 8W

Output: 6V DC 600mA

Connector: 2.1mm jack

The KIWI requires 5V of DC power. The GPS can take anything from 6V to 40V. So, since the power supply outputs 6V, I use a simple voltage divider to get the 5V I need to the KIWI.

Construction:

The KIWI circuit was constructed on a 2½” x 4½” proto-board. Solid core wire was used for most connections. Twisted core wire was used for the GPS RJ-45 connection. The picture below shows the completed circuit:

The circuit is mounted into a plastic enclosure box.

Using the KIWI OSD:

KIWI connected to video in and out and to the GPS:

Initial Setup:

- Turn on the power to the KIWI

- Connect the video input and the video output

- Connect the GPS

- Display:

  - The onscreen display will display “RS232 or 1PPS Absent” at first

  - Then “KIWI OSD V2 FIXWAIT” will be displayed while the GPS attempts to lock onto satellites.

  - The latitude and longitude will then be displayed followed by satellite information

  - UTC date is displayed

  - “KIWI OSD V2 FIXWAIT 1… 2… 3…     … 9”

  - UTC time, end of frame time, start of frame time, counter displayed

This is a picture of the official US time from the NIST webpage with the KIWI onscreen display overlaid:

The bottom of the video screen has this format:

HH:MM:SS EEEE OOOO FFFFFF

where:

HH:MM:SS 	= Hours, Minutes, Seconds UTC time

EEEE     	= End of Frame time (in milliseconds)

OOOO     	= Start of Frame time (in milliseconds)

FFFFFF   	= Contiguous field count since initial GPS sync

Note about the INFO Switch:

Geoff notes the function of the INFO switch on his webpage:

After you have finished the timing run, pressing the "info" switch requests that Kiwi OSD wait for the first valid GPS fix - and then compare its internal clock to the GPS.  If they agree, a message "PREVIOUS TIMES OK" is displayed on screen. If GPS does not agree, "ERROR: USE FIELD COUNT" is displayed, alerting the user that there has been a previous GPS glitch, and to use the video field count to extract timing information.

Pressing the "info" switch after reading the above summary, gives the OSD chip a RESET, and so another timing run is started.

Example of KIWI OSD in action:

I attempted to measure the acceleration due to gravity with the aid of the KIWI. I took a video of a ball dropped from a certain height. The KIWI allowed me to measure precisely when the ball was dropped and when the ball hit the ground. By knowing the free fall time, the gravitational acceleration can be calculated:

Video frames of the ball drop:

The pictures may not be clear to you. Here is the time for each frame:

Frame 1:	19 38 31	439	422	26701

Frame 2:	19 38 31	506	489	26705

Frame 3:	19 38 31	572	556	26709

Frame 4:	19 38 31	639	622	26713

The distance of the fall is 8 inches or 0.2032m (The distance is measured from the top edge of the ball). In frame #1, the ball is stationary throughout the exposure. In frame #2, the ball has begun to fall. Therefore, I will say the start of the fall occurred at the end of frame #1, so at 19:38:31.439

In frame #4, the ball has made contact with the ground. It looks to me to be at the end of the exposure. I will use the end time for frame #4 as the end of the fall, so at 19:38:31.639

The free fall time of the ball is then: 0.639s – 0.439s = 0.200s

The gravitational acceleration is then calculated by:

The uncertainty in the timing of the frames is half the exposure time. So, for frame #1 the start time is 422ms and the end time is 439ms so the uncertainty is 9ms

So my calculation for the uncertainty in g is:

g = (10.16 ± 0.9) m/s2

The accepted value for g is 9.81 m/s2

My value differs with the accepted value by 3.5%

Although I could have gotten closer to the accepted value by having a greater distance for the ball to free fall, my quick measurement and calculation worked fairly well.

Epilogue:

This manual is meant to give a summary of my work in building this device. I tried to explain the function and use of the KIWI OSD as straight-forward and simple as I could. For a much more detailed explanation on the operation of the KIWI OSD, please visit Geoff Hitchcox’s webpages. He is also extremely helpful through email (see page 3).

Special Thanks to

Geoff Hitchcox

Dr. Michael Richmond

Dr. Linda Barton

Bill VanDerveer

# Text boxes

1   2   3   4   5   6   7   8

RJ-45 Connector (head on view)

6.2Vdc

100Ω

10kΩ

KIWI OSD   4.7V

GPS (Pin1)   6.2V

+

-

Ground

KIWI-OSD

GPS

Video In

Video Out

KIWI OSD

NIST time

UTC time

HH:MM:SS

Field Counter

End of Frame time (ms)

Start of Frame time (ms)

Frame 1

Frame 2

Frame 3

Frame 4